Tube Adaption Procedure

Optimus with R5.x BuCo

This procedure MUST be used with R5.x BuCo but will also work on other releases.

Turn Off generator

Remove:	EZX 23 & EZX 43 (Bucky TH)

EZX 42 & EZX 43 (EZ Diagnost)

Turn On generator  (system may take up to 2 min. to get green light with the CAN disconnected)

# Log into Generator with XRG Scope

### Changing Data Set A for tube Adaption

## Bucky TH	Use Registration Device 4 (Free Cassette)

### EZ Diagnost 	Use Registration Device 8 (Free Cassette) for adapting both tubes

Select:		OptimusXRG90/Program/RegistrationDevices/RGDV?/RGDV?DataSetA

#### Note the original settings before changing the parameters in the next step!

Program:	Syncmaster = No

Mounted Radiographical Controller = None

F2 to transmit

Reset Generator

Select:	Registration Device 4 or 8 (Free Cassette) on desk.

### Tube Adaption

Select:	OptimusXRG90/Adjust/TubeAdaption

Select Tube to be adapted

Select Focus  (Adapt small 1st)

F2 to transmit

System is now ready for adaption.  Press and hold hand switch until adaption is complete.

### Restore values in Data Set A

Select:		OptimusXRG90/Program/RegistrationDevices/RGDV?/RGDV?DataSetA

Return the following parameters to the original settings.

Program:	Syncmaster = (Yes)

Mounted Radiographical Controller = (BuCo1)

F2 to transmit

Reset Generator

Power off system

Connect:	EZX 23 & EZX 43 (Bucky TH)

EZX 42 & EZX 43 (EZ Diagnost)