Note: This revised section must be filed into the Customer’s Service Manual to replace the existing Alignment of "Function Unit kV"!

2.		Alignment of "Function Unit kV"

- 2.1		General information

The actual value of the set kV shall be attained at least after 2 ms. At kV rise phase there shall be no kV break-in nor a kV overshoot.

- The "factor for duty cycle" is based on an adapted tube and determines at local mains voltage and mains resistance conditions

- the kV rise phase and

- the kV behaviour during the exposure in falling load technique

as it takes into account the tolerances of the following FRU’s:

PCB EZ 130 																

kV_control_3 = 50kW 		1 converter		4512 108 0908x

kV_control_4 = 65/80kW 	2 converters		4512 108 0910x

a complete power converter unit Q										

kV_power PCB(s) Q100	(part of the power converter unit)

IGBT transistors			(part of the power converter unit)

resonance capacitors			(part of the power converter unit)	

high tension transformer													

An exchange of one of the  marked parts requires a realignment of the "factor for duty cycle"

The "factor for duty cycle" is stored in the memory of PCB CU EZ139. If CU has to be replaced one can reload the CU Complete backup (with the actual factor) to the NVRAM memory or one has to re-align the "factor for duty cycle".

GAIN_IN

X23

PI - Controller

NV HT

UCOMP

X26

AV HT

X3

X6

Ground

Data table Z

CU Z139

Factor Duty Cycle

PWM

Converter

HT Tank

During alignment this "factor for duty cycle" must be entered via XRGSCOPE. The influence of this factor as a correction value for the Z-data table will be monitored as the UCOMP signal, since the PI controller is deactivated by the grounded GAIN_IN signal.

  - Connect and set the scope

Channel 1 = EZ130 X3	AVHT		20kV/V	1V/div		Zero-line at bottom of screen

Channel 2 = EZ130 X26	UCOMP					1V/div		Zero-line 2 div from bottom of screen

Ground = EZ130 X6 	GNDA	all PCB versions

EZ130 X32	GNDA	PCB  ≥ 4512 108 09083     kV_3 (50kW)

PCB  ≥ 4512 108 09104     kV_4 (65/80kW)

Trigger = external (preferred) 	CTRL_X_C/		backpanel EZ X74															\ negative slope

or internal channel 1	AVHT	 EZ 130 X3	/ positive slope at +3V

Time base = 	5 or 10ms/div			trigger delay -1div

Note:	A Digital scope should not have any other ground connection than the ground of the probe. 
A Mains driven scope must be isolated from earth potential, otherwise it might display artefacts.

- 2.3		Deactivate kV controller

- Connect EZ130 X23 GAIN_IN and X6 GNDA with a short link (use a short wire).

Caution: This alignment requires an exposure with high kV!  
     Be sure the tube has been warmed up before.

2.4	Set exposure data

- Notes:

- Apply this adjustment at the tube that has the highest kV limit programmed.

- a) Standard: If the tube limit is 150kV carry out the duty cycle factor adjustment with 141kV; 
b) Special: in case the programmed application limit of the only tube is 125 kV use special setting for this adjustment.

- Any tube arcing during this adjustment requires to execute the tube conditioning next (new procedure is attached). Start over this adjustment if the tube becomes stable.

- Set kV and mA values according to the programmed tube limits:

a) Standard:   141kV		100 mA 	at 	kV_3 	(50kW)

200 mA 	at 	kV_4 	(65/80kW)

b) Special:      125kV		100 mA 	at 	kV_3 	(50kW)

200 mA 	at 	kV_4 	(65/80kW)

- Set exposure time:		40ms

- 2.5	Adjust the "factor for duty cycle" via service software XRGSCOPE by measuring UCOMP with the scope.

- Connect the service PC and start XRGSCOPE:

XRGSCOPE

- SELECT UNITFU-kVADJUSTIGBT Pulse Width Correction Factor for duty cycle:

- Set the starting value "Factor for duty cycle" to 1.00

Note: If the UCOMP value does not match the requirements type in another "Factor for duty cycle" value, <TRANSMIT> the screen and push the active RGDV button to get the new value validated.

- Switch an exposure.
The values are measured in the stationary condition. The transient behaviour at the beginning of the exposure is not taken into account.
Result: In standby the UCOMP value is at about +11V, during exposure the mean value UCOMP shall be as given in table 1a or 1b, refer to picture 1:

picture 1 oscillogram of a perfect alignment for 141 kV					picture 4 UCOMP for 141 kV

- a) Standard setting 141kV:

- Read the mean value of UCOMP for 141kV (see scope picture 2 or 4), correct the factor for duty cycle till UCOMP meets the required reference of + 1Volt.

| kV setpoint | mA setpoint | PCB type | UCOMP | tolerance | KV peak of AVHT | Factor for duty cycle | Date: |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 141 kV | 100 mA | PCB kV_control 3: | + 1V | +1V / - 0,5V | 138 kV |  |  |
| 141 kV | 200 mA | PCB kV_control 4: | + 1V | ± 0.5V | 138 kV |  |  |

Table 1a: duty cycle factor, settings 141kV (150kV limit)

Example how to correct the factor for duty cycle:

- PCB kV_control 3:
If the mean value of UCOMP is

> + 2    V 	increase the factor for duty cycle by increment of 0.01

< + 0.5 V 	decrease the factor for duty cycle in steps of 0.01

- PCB kV_control 4:
If the mean value of UCOMP is

> + 1.5 V increase the factor for duty cycle in steps of 0.01

< + 0.5 V decrease the factor for duty cycle in steps of 0.01

- Check also the kV peak value AVHT (not the overshoot), it shall be 
138kV	 for 141kV setpoint 	(see scope picture 3)

- Remove short link EZ130 X23 GAIN_IN.

- Record the findings in table1a.

picture 2	Optimus 80kW with 141 kV, UCOMP = +1V 					picture 3		HTAV	 without overshoot

---------------------------------------------------------------------------------------------------------------------

- b) Special setting 125 kV:  in case there is no other tube connected that stays for higher kV.

- Read the mean value of UCOMP for 125kV (in principle picture 2 or 4)

- Correct the factor for duty cycle till UCOMP meets the required reference of 0 Volt.

| kV setpoint | mA setpoint | PCB type | UCOMP | tolerance | KV peak of AVHT | Factor for duty cycle | Date: |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 125 kV | 100 mA | PCB kV_control 3: | + 0V | +1V / - 0,5V | 125 kV |  |  |
| 125 kV | 200 mA | PCB kV_control 4: | + 0V | ± 0.5V | 125 kV |  |  |

Table 1b: duty cycle faktor 125 kV limit

Example how to correct the factor for duty cycle:

- PCB kV_control 3:
If the mean value of UCOMP is

> + 1 	  V 	increase  the factor for duty cycle in steps of 0.01

< -  0.5 V 	decrease the factor for duty cycle in steps of 0.01

- PCB kV_control 4:
If the mean value of UCOMP is

> + 0.5 V 	increase  the factor for duty cycle in steps of 0.01

< -  0.5 V 	decrease the factor for duty cycle in steps of 0.01

- Check also the kV peak value AVHT (not the overshoot), it shall be 125kV	for 125kV setpoint.

- Remove short link EZ130 X23 GAIN_IN.

- Record the findings in table 1b.