Optimus Rotor Control replacement with unit

Dual Speed Rotor Control 4512 104 71421/61

Before replacing a high speed rotor control unit carry out the following measures:

If breakers ENF1 / 2 / 3 tripped due to an unknown reason they should not be turned back on. First check:

1)		mains supply and distribution

-	cut the generator from mains:

-	remove mains lines from the wall junction box (Optimus RAD = Bucky)

or

-	turn off the room mains switch (Optimus R/F = Fluoro + Bucky)

-	check at open breakers (off condition) that no contact is welded (Ohm-meter)

ENF1		L1 - T1						Optimus 50			2422 129 15514

L2 - T2						Optimus 65/80		2422 129 15513

L3 - T3

ENF2		L1 - T1						Optimus 50/65/80	2412 129 00341

L2 - T2

L3 - T3

ENF3		L1 - T1						Optimus 50/65/80	2422 129 15665

L2 - T2

L3 - T3

-	check that no contact of contactor ENK2 is welded (mechanical check: if snap-on auxilary 	contacts can not be removed from the main contactor at least one contact is welded)

ENK2	phase 1		1 - 2			Optimus 50/65/80			2422 132 06621

phase 2		3 - 4			3 auxilary contacts NO		2422 135 04355

phase 3		5 - 6			1 auxilary contact NC		2422 135 04023

!!	In case one of the contacts is welded the converter DC discharge time will be

!!	much longer as the discharging auxilary contact ENK2 41-42 can not be closed.

!!	(600V DC discharge time to < 60V DC approximately 40 seconds,

!!	 with open contact  > 11 minutes !! to be < 60V DC)

2)		measurement of stator resistance

The defect of the rotor control unit might have been caused by the stator coils.

Measure the stator phase resistance (for all tubes connected)

at the tube (disconnect stator cable, but all four links must be present and fixed)

phase U-V		101/103 - 107/109	normally 11 Ohms ± 1 Ohm

phase V-W		208/210 - 202/204	normally   9 Ohms ± 1 Ohm

If measured at the stator relays EWGX11/12 or the stator cable terminal EX1100 add the cable resistance:

cable resistance 	(10m single length) + 0.27 Ohms

(3 x 1.31mm2	)		(15m single length) + 0.40 Ohms

(20m single length) + 0.53 Ohms

3)		measurement of stator inductance (if measuring tool available)

It is not sufficient to measure the stator phase resistance only. The resistance measurement might pretend the stator coil is ok. A short-circuit of windings can not be found with an Ohm-meter.

With an inductance meter measure

at the tube (disconnect stator cable, all four links must be present and fixed

phase U-V		101/103 - 107/109		normally		57mH ± 10 %

phase V-W		208/210 - 202/204		normally  	34mH ± 10 %

If the deviation is > ± 10 % the tube (stator) must be replaced to prevent from damages of the rotor control PCB Y100.

4)		first turn-on after rotor control ‘event’

If chapters 1)...3) are error-free

or

if parts of the mains distribution have been replaced:

Establish the mains cable at the wall junction box or turn the room mains switch back on.

Turn on all breakers except ENF3 !

Reason: If the generator had been in preparation condition and had been cut from mains during a fuse trip the CPU memorizes that it did not send a brake command.

It will do so during the next turn-on.

Once the generator is back on again the command will be send. CU does not get and does not wait for a brake confirmation of RoCo.

To be on the safe side and to prevent from further damage the brake command should take place without any power for the RoCo drivers, therefore ENF3 should be off.

The generator will display an error message 10LU = intermediate DC out of range (on the desk and/or on the service PC if online) due to the missing power supply.

If the error does not clear by itself push Reset button on the control desk.

If the generator comes to a standby ready state (check ready conditions of the system if not) proceed to the tube extension chapter 5).

5)		tube extension EWG check (if not present proceed to 6))

Breaker ENF3 should still be off but the generator be on.

Check that one of the stator relays EWGK11/12 is energized (small brown indication pin not at left edge of the slit). The contactor must not be activated manually.

The contact resistance can not be measured with a regular Ohm-meter as the small current of the meter pretends a very high resistance (as a matter of the silver plated contact area).

It must be measured invasive with at least 2 A passing the closed contact. As the contact  should have a typical resistance of < 0.1 Ohm, the voltage measured via the contact should not increase 0.2 Volt (with 2A).

EWGK11	phase U		1 - 2

phase V		3 - 4

phase W	5 - 6

EWGK12	phase U		1 - 2

phase V		3 - 4

phase W	5 - 6

If the measured value is higher, exchange the contactor (2422 135 03942).

6)		replacement of rotor control unit

Switch off the generator with ENF1 or with the room mains switch.

If chapters 1)...5) are error-free or parts have been replaced, exchange rotor control unit (4512 104 71421 for release 3 generators, 4512 104 71461 for release 2 generators).

Chokes which have been added to the previous RoCo unit must not be mounted to the new unit, as this choke (looking different) already exists on the RoCo rectifier PCB Y200.

Additional NTCs of power-on limitation kits 4512 104 91281 (FCO 00 135 007) must also be discarded.

The new units 4512 104 71421/461 are already equipped with firmware 4512 113 22322.

Release 2 generators:

It is sufficient to replace the unit.

Turn on ENF3. Turn on the generator and check all functions.

Release 3 generators:

Turn on ENF3. Turn the generator on.

The generator might come up with error messages 10LU and 10TD.

Reload the tube data from file TUBE_R3.TDL (file size 176.996B, date 15.04.99), previous versions don't work.

Afterwards adaptation has be carried out.